## Slide 1: INTRODUCTION TO AUTOMATIC FILM PROCESSING: EQUIPMENT AND TECHNIQUES
RAD. PATRICK COLLINS
MXT205

## Slide 2
Automatic film processing uses specialized machinery to quickly and uniformly convert a latent image on exposed film into a dry , visible radiograph or photograph.
KEY EQUIPMENT
Film feed section: The entry slot where the film is inserted, which activates the transport motor and replenishment system via a microswitch.
Roller transport system : A series of racks, crossover turn-arounds, and guide shoes that physically pull and push the film through each chemical tank at a constant speed.

## Slide 3
Processing tanks: Compartments holding temperature – controlled solutions for developing , fixing , and washing:
Developer Tank: converts exposed silver halide crystals into black metallic silver ( typically maintained around 33 C / 96 F ).
Fixer Tank : Dissolves unexposed / underdeveloped silver halide and hardens the gelatin emulsion.
Wash Tank: Uses circulating water to flush away residual processing chemicals.
Recirculation and filtration system: continuously agitates, mixes , filters , and maintains even temperatures across the chemical baths.

## Slide 4
Replenishment system: Automatically pumps fresh chemical solution ( or replenisher ) into the developer and fixer tanks based on the size or running time of the processed film to maintain chemical activity.
Drying system: uses heated air ( around 50 C to 55C / 120F to 130F ) blown across the film surface via squeegee rollers to yield a completely dry film upon exit.

## Slide 5
STANDARD TECHNIQUES AND WORKFLOW
Total cycle Time : Modern automated medical or industrial processors complete the entire dry – to –drop sequence in roughly 50 seconds ( compared to an hour or more in manual processing ):
Developing – 22 seconds
Fixing – 22 seconds
Washing – 20 - 26 seconds
Drying – 20 seconds

## Slide 6
Feeding Technique : Films are fed straight into the entry tray in a staggered pattern ( rather than overlapping side – by –side ) to prevent roller jams , overlapping , or chemical contamination .
Standby mode: When idle, the machine drops into a low – power standby state where the roller drive stops or slows down to conserve water and electricity until a new film triggers the sensor.

## Slide 7
DAILY AND MONTHLY MAINTENANCE
This requires routine cleaning of crossover rollers to prevent “” pi lines “ or guide – shoe marks, alongside periodic tank draining , rack scrubbing, and sensitometric testing
MEDICAL X-RAY AUTOMATIC PROCESSING
A medical x-ray automatic processor is an electrochemical device designed to automate and standardize the chemical processing of radiographic films. It converts a latent ( invisible ) image on an exposed x-ray film into a permanent , manifest ( visible ) radiograph in a matter of seconds.

## Slide 8
1. PRINCIPLES OF AUTOMATIC PROCESSING:
The core principle of automatic processing is standardization of time , temperature , and chemical activity. It operates on the same chemical sequence as manual processing – developing, fixing , washing, and drying – but compresses the entire timeline ( typically down to 90 seconds ) by utilizing:
High chemical concentration: Formulated specifically to handle aggressive development speed.
Elevated temperature: Processing solution are heated to maximize chemical activity.

## Slide 9
Continuous mechanical agitation: Constant contact between the rollers and the film emulsion forces fresh chemicals into the film surface while squeezing out exhausted solutions.
2. CONSTRUCTION AND DESIGN MECHANISM:
The processor consists of several independent but integrated sub-system:
Transport system: A network of gears, chains, a drive motor, and precision roller racks designed to pull the film along a serpentine path through the liquid tanks and the drying chamber at a fixed speed.

## Slide 10
Temperature control system: Uses immersion heaters, thermostats, and heat exchangers to precisely regulate and maintain the developer solution at 33C to 35C ( 92C to 96C ).
Circulation / recirculation system: Continuously pumps and agitates the developer and fixer solution to prevent thermal stratification, filter out gelatin particles, and maintain chemical homogeneity .

## Slide 11
Water system: Supplies fresh , flowing water to rinse chemical residues from the film, preventing archival degradation ( hypo-retention strains ).
Dryer system: Uses a high – velocity blower, heating elements ( 40C to 65C ), and air tubs to evaporate remaining surface moisture from the film.
[ feed tray ] > [ developer Tank ] > [ Fixer Tank] > [ Wash Tank ] > [ Dryer Chamber ] > [ Receiving Bin ].

## Slide 12
3. FEED TRAY AND FEEDTRAY TYPES
The feed tray is the entry platform where the exposed radiographic film is introduced into the unit.
Mechanism: It contains micro switches or infrared optical sensors that detect the physical presence, size ,and orientation of the inserted film. This detection signal activates the drive motor , the dryer, and calculates the required volume of replenisher chemicals.

## Slide 13
Feed tray types:
Darkroom feed tray: A standard tray located completely inside a darkroom under safelight condition where the operator manually feeds exposed films extracted from cassettes.
Daylight loader ( daylight feed tray ) : A light – tight, enclosed chamber attachment fitted with sleeves or an automated cassette – unloading mechanism. This setup allows the operator to insert their hands and feed the film in normal room light without exposing the film emulsion.

## Slide 14
4. ROLLERS ( TRANSPORT COMPONENTS )
The mechanical movement relies heavily on specialized rollers clustered into functional racks:
Entrance Rollers: Located immediately past the feed tray, they grab the film and trigger the processing cycle.
Transport / Deep- Tank Rollers: Positioned inside the liquid tanks, these planetary rollers guide the film downward and upward along the rack assembly.

## Slide 15
Cross- over Rollers: Situated above the solution levels between adjacent tanks ( e.g., developer to fixer). They feature miniature guide shoes to redirect the film from one path to the next.
Squeegee Rollers: Located at the exit of the wash tank. They apply physical pressure to strip excess water from both sides of the film emulsion before it enters the drying chamber, reducing the energy needed for thermal drying.

## Slide 16
5. REPLENISHMENT SYSTEM
As films pass through the chemistry ,solutions wear out ( developer oxidizes and loses electron. Fixers accumulates silver ions and dilutes ). The replenishment system injects exact doses of fresh chemicals from storage reservoirs into the active processor tanks to maintain stable chemical potency.
Volume ( pulse ) replenishment : Triggered dynamically whenever a film crosses the entry sensors . The pump inject a specified amount based on the length or surface area of the processed film.

## Slide 17
Flood ( timed ) Replenishment: Configured for low- volume clinics . The processor automatically pulses fresh chemistry at a fixed time interval ( regardless of film throughout ) the combat natural aerial oxidation.
FORMULATING CHEMICALS FOR AUTOMIC PROCESSORS
Because automatic processing occurs at high speed and elevated temperatures , standard manual chemistry cannot be used. The chemistry must contain specialized hardening compounds to protect the film from swelling or melting.

## Slide 18
S/N | SOLUTION COMPONENT | ACTIVE CHEMICAL AGENTS | PRIMARY FUNCTION
1 | Developer | Hydroquinone and phenidone Sodium sulfite, potassium bromide Glutaraldehyde | Converts exposed silver crystals to black metallic silver , preservative against Aerial oxidation, hardiner –( prevents excessive emulsion swelling) Restrainer ( pre vents unexposed crystals from developing )
2 | Fixer | Ammonium thiosulfate, Acetic Acid, Potassium, Alum, sodium sulfite | cleaving agent ( dissolves unexposed silver halide crystals ). Acidifieis/buffer( neutralizes residual alkaline developer). Hardener( shrinks and permanently hardens the gelatine). Preservative.

## Slide 19
ROUTING OPERATING PROTOCOLS
A. STARTING ROUTINE
Check the chemical volumes in the external replenishment tanks
open the main water supply valve to initiate flow into the wash tank
Ensure the active tank drain valve are tightly closed.
Turn on the main electrical power toggle
Allow 15 – 20 minutes for the chemical heaters to reach stable operating temperature ( indicated by a ready light ).

## Slide 20
vi. feed a clean “cleaner film” ( dummy film ) through the processor to lift any dried chemical crust 0r debris off the roller pathways before processing patient films.
B. SHUTDOWN ROUTINE
Turn off the main electrical power button
Turn off the main water valve to prevent waste or flooding.
Remove the top processor splash cover to reduce condensation and prevent chemical crystals from dripping back into opening tank
Leave the main top lid slightly propped open to vent residual humidity overnight.

## Slide 21
S/N | PROBLEM/ ARTIFACT | LIKELY CAUSE | SOLUTION
1 | Damp or wet film | Dryer temperature too Low , exhausted fixer, Hardener; high room Humidity | check dryer heating element; verify fixer strength
2 | Guide shoe marks/ Scratches | Misaligned or dirty roller guide shoes scratching The emulsion | clean the racks, adjust guide shoe alignment
3 | Pi lines ( 3.14-inches Repeats) | Dirt, chemical crust, or a physical dent on a specific rollers | clean thoroughly, replace damaged rollers
4 | Chemical fog/dark film | developer temperature too High, over replenishment | calibrate thermostat Down, adjust replenishment rates
5 | Film jamming | misaligned rollers; feeding Films too close together or Crookedly. | stagger film feeds; Inspect chain alignment.
8. TROUBLESHOOTING GUIDELINES

## Slide 22
9. QUALITY ASSURANCE ( Q A )
A. DAILY Q A PROTOCOLS:
Sensitometry and Densitometry: Expose a control film using a sensitometry, process it , and measure the optical density steps ( Base + fog , speed index , contrast index ) using a densitometer to confirm chemical stability.
Temperature verification: manually check that the developer temperature matches the digital readout.
Crossover Roller Cleaning: Wipe down the crossover assembles first thing in the morning.

## Slide 23
B. WEEKLY Q A PROTOCOLS
Replenishment Rate Verification: Measure the actual volume of developer and fixer pumped per 14* 17 inch film to ensure accurate delivery.
Deep Clean Racks: Remove the main developer and fixer racks , rinse them with warm water , and scrub away any gelatin build up or chemical scales.
Water Filter Inspection: Inspect and replace the water supply system filter if a drop in water pressure or flow rate is observed.

## Slide 24
THANK YOU